## Slide 1
Metrology calibration · Measuring tool management · MSA
From ledger traceability to measurement system trustworthiness · Series of integrated practical training
If your eyes are not accurate, all quality data is built on sand · Knowledge No. 6.2.1
QTank
Knowledge No. 6.2.1 · MSA and metrological traceability
qtank.org

## Slide 2
Training Objectives
Supporting article id=225 / 22 / 206 / 70 · It is recommended to train within 3 to 4 hours
Establish a closed loop of accounting, marking and calibration plans for the entire life cycle of measuring tools
Distinguish between calibration and verification, be able to read certificates and conduct traceability assessment of non-conformities
Explain the measurement traceability chain and the "one-third" uncertainty principle
Master the five characteristics of MSA and GR&R sampling, execution, judgment and improvement
Can do alternative strategies for bias/linearity/stability and destructive measurements
Able to draw the minimum path for digital implementation of "Calibration Management + MSA"
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 3
Course Map and Article Mapping
module | theme | Main material
01 | Why measurement system is the foundation of quality | 225 Introduction
02 | Measuring tool classification·Ledger·Identification | 225 § 2
03 | Calibration plan·certificate·disqualification | 225 § Three
04 | Measurement traceability and uncertainty | 225 § Four + 70
05 | On-site use and maintenance and misunderstandings | 225 § 5
06 | MSA Five Characteristics and GR&R Practice | 22 + 206
07 | Bias/Linearity/Stability·Destructiveness | 206
08 | Digitalization and sustainable mechanisms | 225 + 206
09 | Workshop and Implementation Checklist | comprehensive
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 4
01
Beginning: Invisible Foundation
The ruler is inaccurate, SPC/CPK are all illusions

## Slide 5
The most common "minor issues" in audits
❌ On-site typical
The expired caliper is still in use in the drawer The certificate expired for three months and was left unattended The ledger does not match the physical object No calibration mark/color code confusion on site
⚠️Real consequences
Accidental killing of qualified products → waste Let go of unqualified products → Customer complaints/recalls Industry survey: >40% of companies have systemic problems ISO 7.1.5 / IATF MSA+ traceability mandatory
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 6
Position this training in one sentence
Measurement calibration and measuring tool management: ensuring that the "ruler itself" is traceable within the validity period
MSA: Prove that the measurement system composed of "employers + measuring tools + methods" is reliable enough for decision-making
Both are indispensable: calibration qualified ≠ MSA qualified; MSA passed ≠ lifetime maintenance-free
Goal: Data credibility → Determine credibility → Improve credibility
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 7
02
Measuring tool classification and ledger
Take care of your property first, then talk about accuracy

## Slide 8
Three sets of classifications: Accuracy/Object/Management Level
Dimensions | Typical category | Management points
Accuracy | Standard instruments / work measuring instruments / monitoring instruments | The standard is at the top of the traceability chain
object | Length·Mechanics·Temperature·Electricity·Special inspection tools | Assign people and environment according to professional abilities
grade | A Mandatory verification/B Enterprise standard/C General | A/B strictly controls cycle and qualifications
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 9
The core fields of the "account book" of the measuring tool ledger
Unique number (type-department-serial) + name specification + manufacturer/factory number
Accuracy/gradation value, using department/storage location, responsible person
Calibration cycle, latest calibration date/validity date, external school/internal school/exemption from school
Status: normal / disabled / scrapped; the certificate number can be checked back
Iron rule: Accounts should be consistent; it is recommended to take physical inventory every quarter
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 10
On-site visual identification (color code)
color | meaning | action
green | Qualified and within validity period | Can be used normally
yellow | Limited time/limited use | Use with restrictions
red | Unqualified or disabled | Isolate, prohibit measurement and release
scrapped | Log out permanently | Physically destroyed or explicitly sealed
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 11
03
Measurement calibration system
Complete link from program to certificate

## Slide 12
Calibration vs Verification
Calibration Calibration
Determine the relationship between indication value and standard value Issue calibration certificate/indication error Generally unable to enforce legal effect Enterprises build themselves around "data trustworthiness"
Verification
Mandatory comprehensive assessment by statutory bodies Issue a calibration certificate Trade settlement/security/medical/environment, etc. The scope is determined by legislation
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 13
How to make a calibration plan before it can be executed
Period: type × frequency of use × environment × stability × historical data comprehensive determination
Reference: Caliper is about 12 months; micrometer is about 6 to 12 months; torque wrench is about 3 to 6 months
New purchase: first calibration upon arrival to verify factory accuracy
Method: external school (CNAS/statutory institution) or internal school (enterprise standard + certified personnel)
Execution: Rolling monthly plan for centralized school delivery, leaving a buffer, and prohibiting "cramming" on due dates
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 14
Calibration result judgment and processing
Qualified renewal
Adjustable→recalibration
Not repairable→disabled
Impact assessment traceability
Update ledger labels
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 15
"Impact assessment" after failure - the most easily missed items
ISO/IATF: Measurement results from the last qualified calibration to the current failure must be evaluated
Scope: Involves batches, release decisions, customer deliveries, special characteristics
Output: continue to use/strengthen re-inspection/quarantine recall/notify customers (according to CSR)
Evidence: Assessment records and disposal closed-loop archiving for review and customer complaint tracing
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 16
Four things about calibration certificate management
Certificate number ↔ measuring tool number one-to-one correspondence
Validity period warning (system/board notice in advance)
Electronic/paper files are created based on measuring tools to retain historical trends.
Receipt review: standard instrument information, environment, uncertainty, and results are complete
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 17
04
Measurement traceability and uncertainty
The ruler must be connected to the national datum

## Slide 18
Traceability pyramid (top to bottom)
national benchmark
social public standards
The highest standards of the enterprise
work measuring instrument
Product tested
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 19
Key points of traceability management practice
External schools choose CNAS accredited laboratories or legal metrology institutions
Internal school: establishing corporate standards, training metrologists, and regular external traceability of the highest standards
No traceability = self-talk: customer re-inspection disputes often arise from inconsistent standards
The uncertainty on the certificate must be read: otherwise "passed" may not be enough
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 20
Uncertainty and Laboratory Comparisons (Article 70)
Measurement uncertainty
Any measurement result has an interval Consider the “grey zone” in quality judgment Rule of One-Thirds: Calibration U ≤ 1/3 of the allowable error
Laboratory comparison
Multiple laboratories testing the same sample Verify capability and consistency Uncertainty statements are verified by comparison 17025 / Commonly used in customer-specified laboratories
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 21
05
On-site use and maintenance
Passed on calibration day ≠ Still accurate today

## Slide 22
Confirm 10 seconds before use (write SOP)
Is the calibration label within the validity period?
Appearance: bumps, deformation, rust
Zero position/clean measuring surface (caliper, micrometer)
Digital display: power and display are normal
Habitual actions can eliminate most risks of misdetection
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 23
Daily maintenance points (excerpt)
type | key action
General measuring tools | Wipe clean, rust-proof, special box and cabinet, remove the battery for a long time
Special inspection tools | Clean cutting fluid/iron filings, check wear, and maintain constant temperature and humidity for precision parts
Standardizer | Clean with absolute ethanol, wear gloves, and do not touch the measuring surface with bare hands
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 24
Three major on-site misunderstandings
Misunderstanding ① "It's accurate once it's calibrated" → Periodic verification and zero position check before use are required
Misunderstanding ② "The more precise the better" → Adequacy principle to prevent over-provisioning and environmental sensitivity
Misunderstanding ③ "It's just a matter for the measurer" → Division of measurement/supervisor/operator
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 25
06
MSA and GR&R
Demonstrate that the measurement system is reliable enough for decision making

## Slide 26
MSA Core Concepts (Article 22 / 206)
Measurement system = measuring tool + person + method + fixture + environment
If measurement variation swamps process variation, SPC/CPK/judgment will be completely distorted
IATF / AIAG MSA 4th: Key characteristics require MSA evidence
MSA is an ongoing mechanism, not a "one-time PPT pass"
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 27
Five properties of measurement system
Dimensions | focus | Typical method
Bias | Accurate or not (relative true value) | Repeated testing of standard parts + t test
Linearity | Does the bias change with the range? | Multiple standard parts regression
Stability Stability | Will it fade over time? | Standard parts control drawing
Repeatability EV | The same person and the same thing fluctuate | GR&R decomposition
Reproducibility AV | Differences between people | GR&R decomposition
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 28
GR&R Sampling: The starting point determines the end point
Samples must cover the full tolerance range, and the production line should not "just pick up 10 pieces" and squeeze them in the center
When the process capability is very high: manual selection or production of samples with upper and lower limits
Recommended: key safety features 10×3×3; general parameters can be simplified to 5×2×2
Measuring tools must be within the calibration validity period; record sheets and blind test disciplines must be prepared in advance
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 29
Implement three disciplines (most prone to errors)
① The measurement order is randomized, and consecutive measurements in order of size are prohibited.
② Insert normal operations between two measurements to simulate the real inspection rhythm
③ The results are kept confidential to the operator and prevent reading inertia and human correction.
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 30
AIAG %GR&R Determination (Don’t One Size Fits All)
%GR&R | determination | action
≤ 10% | acceptable | maintain and monitor
10%～30% | Conditionally accepted | Make decisions based on risk and cost of improvement
> 30% | Unacceptable | Must be improved before use for release/SPC
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 31
%GR&R/TV vs %GR&R/Tol
Relative process variation TV
Look at the ratio of measurement variation to the total process fluctuation More suitable for SPC/process monitoring scenarios When the process is very stable, TV is prone to "falsely high"
Relative tolerance Tol
Look at the ratio of measurement variation to tolerance bandwidth More suitable for conformity determination/inspection scenarios It is recommended to look at both indicators at the same time
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 32
GR&R failed: troubleshooting order
points EV/AV
Check sample coverage
Check method/fixture
Check the ambient temperature
Let’s talk about changing measuring tools again
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 33
"Lower is better" is also a misunderstanding
%GR&R well below 5% may be over-provisioned (±5 mm roughing measured with CMM)
High-precision equipment is more sensitive to the environment and may introduce new variations and higher costs
Principle: Select an “adequate” measurement system based on tolerance and process variation
Also look at ndc (number of distinguishable categories): too low indicates insufficient ability to resolve process variability.
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 34
07
Bias·Linearity·Stability·Destructiveness
Beyond precision, there’s accuracy

## Slide 35
Bias / Linearity / Stability in one sentence
Bias: How much the mean deviates from the reference value (fixed systematic error)
Linearity: Does the bias drift as size increases (full-scale consistency)
Stability: Is the control chart of standard parts out of control over weeks to months?
Calibration "point qualified" ≠ stable during the use period; check during the period to add the timeline
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 36
MSA replacement for destructive measurements
Strategy | Applicable | Main points
Nesting method | Material/adjacent sample | Separating batch variation from measurement variation
Attribute consistency Kappa | Pass/fail determination | Kappa≥0.75 is usually better
Blind test of verification parts | Special damage test | Ability to perform regular spot checks on known results
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 37
08
Digitalization and sustainable mechanisms
From compliance checks to daily management

## Slide 38
Core functions of measuring tool digitization (225)
Electronic ledger + barcode/QR code full life cycle
Automatic early warning of calibration expiration (email/WeChat/system)
Electronic archiving of certificates and trends in health records
Scan the code on site to intercept expired/red label measuring tools
Qualification rate, fault type, frequency of use reports
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 39
MSA Digital Trends (206)
Automatic collection + online calculation GR&R/bias, abnormality warning
Multivariate Measurement (CMM/Imaging) Overall Evaluation
Vision/AI inspection: confusion matrix, ROC, not just traditional GR&R
Goal: Annual review materials → Daily measurement health board
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 40
Suggested paths for digital evolution
Excel+Alert
Scan code ledger
Certificate electronic file
Interface with QMS/MES
Online MSA
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 41
09
Workshop and implementation
See you in two weeks

## Slide 42
Suggested agenda (3-4 hours)
time period | content | output
30′ | Foundation and review pain points | Consensus: Why do it?
45′ | Ledger/Identification/Calibration Plan | First draft of field list
40′ | Traceability·Certificate·Disqualification Traceability | Disposal Checklist
50′ | MSA Five Elements and GR&R Exercise | a characteristic sampling plan
30′ | digital minimum path | 90 day roadmap
25′ | Workshop summary and division of labor | Two-week action items
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 43
Workshop: Choose a "Pilot Feature"
Select 1 critical dimension or quantifiable characteristic such as torque/seal
List: Gage number, calibration status, operator, method, environment
Design GR&R: number of parts × number of people × times + randomization rules
Synchronous check: Whether the ledger fields are complete and certificate U meets 1/3
Output: Complete data collection within one week, and issue judgments and improvement items within two weeks
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 44
Two-week landing checklist
□ Complete the inventory of measuring tools and unify color standards in the pilot area
□ Release rolling monthly calibration plan + person responsible for expiry warning
□ Complete the "Impact Assessment" form and process for unqualified calibration
□ Complete 1 key feature GR&R (including EV/AV interpretation)
□ Decide on the first step of digitization (scan the ledger or electronic certificate file)
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 45
Common pitfalls in auditing and management (general review)
There is a certificate but no ledger / there is a ledger but no physical object
There will be no interception for overdue use; there will be no batch impact assessment for unqualified use.
Only GR&R is done without bias stability; cross-over is done once a year
Samples do not cover tolerances; results are not blindly tested; GR&R is too low, blind pride
Measurement island: not linked to CP, SPC, and release systems
Metrology Calibration·Measuring Tool Management·MSA Series Integrated Training

## Slide 46
Q & A
Supporting think tank articles: id=225 Measurement calibration and measuring tool management id=22 MSA Practical Guide id=206 MSA in-depth practice (GR&R→system) id=70 Measurement uncertainty compared with laboratory Knowledge No. 6.2.1 · This PPT can be downloaded from the database qtank.org

## Slide 47
thanks for listening
qtank.org
QTank
qtank.org